Contoso Electronics

Equipment Reference & Maintenance Guide

Environmental Test Chamber

Contoso ClimaTest EC-2000

Asset Identification

| Asset ID | CE-ETC-3100 |
| --- | --- |
| Category | Reliability Lab |
| Manufacturer | Contoso Environmental Systems |
| Location | Reliability Lab - Chamber 1 |
| Installation Date | 2022-09-15 |
| Warranty Expiry | 2027-09-15 |
| Support Contact | reliability.support@contoso.com ¦ ext. 4710 |

Overview

The Contoso ClimaTest EC-2000 is a temperature and humidity test chamber used for environmental stress screening, thermal cycling, and reliability qualification of electronic assemblies to industry standards.

Technical Specifications

| Parameter | Value |
| --- | --- |
| Temperature Range | -70 °C to +180 °C |
| Humidity Range | 10% to 98% RH |
| Chamber Volume | 225 liters |
| Temp Ramp Rate | 5 °C/min |
| Uniformity | ±0.5 °C |
| Controller | Programmable, 100 profiles |
| Power | 3-phase 400 VAC, 12 kW |

Preventive Maintenance Schedule

| Interval | Task |
| --- | --- |
| Weekly | Check water reservoir and drain; inspect door seal. |
| Monthly | Clean condenser coils; verify temp/humidity accuracy. |
| Quarterly | Inspect refrigeration system pressures. |
| Annual | Full calibration; replace humidity sensor if drifting. |

Safety Guidelines

Chamber surfaces can be extremely hot or cold - use PPE.

Do not open the door during extreme-temperature operation.

Ensure adequate ventilation for refrigerant safety.

Never place flammable or sealed pressurized items inside.

Common Issues & Troubleshooting

| Symptom | Recommended Action |
| --- | --- |
| Cannot reach low temp | Check refrigerant charge and condenser airflow; verify door seal. |
| Humidity error | Refill water reservoir; clean wick; verify sensor. |
| Temp overshoot | Retune PID; reduce load mass; check sensor placement. |
| Frost buildup | Run defrost cycle; inspect for door leaks. |

Contoso Electronics internal maintenance documentation. For service requests, contact reliability.support@contoso.com | ext. 4710 or submit a work order through the maintenance portal.