Contoso Electronics

Equipment Reference & Maintenance Guide

Programmable DC Power Supply

Contoso PowerBench PS-3060

Asset Identification

| Asset ID | CE-PSU-1400 |
| --- | --- |
| Category | Test & Measurement |
| Manufacturer | Contoso Instruments Division |
| Location | Electrical Test Lab - Bench 2 |
| Installation Date | 2023-04-11 |
| Warranty Expiry | 2027-04-11 |
| Support Contact | test.equipment.support@contoso.com ¦ ext. 4210 |

Overview

The Contoso PowerBench PS-3060 is a programmable triple-output DC power supply used to power boards under test with precise voltage and current control, sequencing, and protection features for safe, repeatable bench and production testing.

Technical Specifications

| Parameter | Value |
| --- | --- |
| Outputs | 3 (2 x 0-30 V/5 A, 1 x 0-6 V/3 A) |
| Line/Load Regulation | < 0.01% + 3 mV |
| Readback Accuracy | 0.05% |
| Protection | OVP, OCP, OTP |
| Interface | USB, LAN, GPIB |
| Programming | SCPI command set |
| Power | 100-240 VAC, 50/60 Hz |

Preventive Maintenance Schedule

| Interval | Task |
| --- | --- |
| Monthly | Verify output voltage/current accuracy. |
| Quarterly | Clean intake filter; check terminal tightness. |
| Annual | Full calibration; verify protection trip points. |
| As needed | Inspect binding posts and leads for damage. |

Safety Guidelines

Outputs can source high current - avoid short circuits.

Verify polarity before connecting the DUT.

Do not exceed rated wire gauge for output current.

Enable OVP/OCP to protect sensitive devices under test.

Common Issues & Troubleshooting

| Symptom | Recommended Action |
| --- | --- |
| Output shuts off | OVP/OCP tripped - check DUT and setpoints; reset protection. |
| Voltage not reaching setpoint | In constant-current mode - check load current draw. |
| Inaccurate readback | Recalibrate; verify remote sense connections. |
| No remote control | Check interface cabling and SCPI address configuration. |

Contoso Electronics internal maintenance documentation. For service requests, contact test.equipment.support@contoso.com | ext. 4210 or submit a work order through the maintenance portal.