Contoso Electronics

Equipment Reference & Maintenance Guide

Soldering Station

Contoso ThermaPro SS-90D

Asset Identification

| Asset ID | CE-SOL-0450 |
| --- | --- |
| Category | Assembly & Rework |
| Manufacturer | Contoso Assembly Tools |
| Location | SMT Rework Cell - Station 7 |
| Installation Date | 2024-01-09 |
| Warranty Expiry | 2026-01-09 |
| Support Contact | assembly.support@contoso.com ¦ ext. 4155 |

Overview

The Contoso ThermaPro SS-90D is a temperature-controlled soldering station designed for precision hand soldering and rework of surface-mount and through-hole components. It features rapid tip heat-up, closed-loop temperature control, and ESD-safe construction.

Technical Specifications

| Parameter | Value |
| --- | --- |
| Power Output | 90 W |
| Temperature Range | 100 °C - 450 °C |
| Temperature Stability | ±2 °C |
| Heat-up Time | < 10 seconds to 350 °C |
| Tip Type | Contoso T12-series cartridge |
| ESD Rating | ESD-safe (< 2 Ω to ground) |
| Power | 120/230 VAC, 50/60 Hz |

Preventive Maintenance Schedule

| Interval | Task |
| --- | --- |
| Each shift | Clean tip with brass sponge; re-tin before storage. |
| Weekly | Inspect tip for pitting; check ground continuity. |
| Monthly | Verify temperature accuracy with tip thermometer. |
| As needed | Replace worn cartridge tips and heating element. |

Safety Guidelines

Tip temperatures exceed 400 °C - never touch the tip.

Return the iron to its stand when not in use.

Work in a ventilated area or use fume extraction.

Wear ESD wrist strap when handling sensitive components.

Common Issues & Troubleshooting

| Symptom | Recommended Action |
| --- | --- |
| Tip not heating | Check cartridge seating and station fuse; test with spare tip. |
| Poor solder wetting | Clean and re-tin tip; verify temperature setpoint and flux. |
| Temperature drift | Recalibrate; inspect for damaged thermocouple. |
| Error E-04 on display | Heating element open circuit - replace cartridge. |

Contoso Electronics internal maintenance documentation. For service requests, contact assembly.support@contoso.com | ext. 4155 or submit a work order through the maintenance portal.